FRx specs

Specifications

Defibrillator must deliver therapy using a biphasic truncated exponential waveform and automatically adjust parameters as a function of chest impedance during delivery of the waveform.

Defibrillator must be available for use on any patient of any age, including small children and infants.

Defibrillator must deliver 150 Joules nominal energy to a 50-ohm load.

Defibrillator must be able to deliver a shock within 8 seconds after the end of the CPR pause

The defibrillator must fully disarm the capacitor internally under any of the following conditions: - A no shock decision is reached. - The defibrillator is turned off. - 30 seconds after arming if the shock button is not pressed. - The defibrillation pads are removed from the patient.

The defibrillator must provide natural sounding, high fidelity voice prompts.

The defibrillator must provide voice prompts that are responsive to/interactive with the user’s actions with the adhesive pads

The defibrillator must possess the ability to do CPR voice coaching

The defibrillator must remind the user to call Emergency Medical Services

The defibrillator patient analysis algorithm must detect electrical noise (artifact) that could interfere with the device’s ability to perform analysis

The defibrillator must reject pacemaker artifact

The defibrillator must analyze the heart rhythm without requiring the user to initiate analysis.

Defibrillator Size, Weight and Durability Specifications:

Specifications

Defibrillator unit must weigh no more than 3.5 pounds with the battery installed and pads attached.

The defibrillator must withstand a drop of 1.22 meter (4 feet) to any edge, corner, or surface and remain operational.

The defibrillator must be no larger than the following dimensions: (7.1” H x 8.9” W x 2.4” D)

The defibrillator must meet standards for use in aircraft and within altitude of -1312 to 15,000 ft.

The defibrillator must be able to withstand a 500 kg (1100 lbs) crush test

Must have IP55 rating

Defibrillator Maintenance and Testing Specifications:

Specifications

The defibrillator must conduct an automated self-test at the following intervals/events: - Daily - Weekly - Monthly - Following the Insertion of battery (BIT test) - When the defibrillator has been stored outside prescribed temperature range - After attaching pads

- The defibrillator must provide an active visual and audible indicator that it has passed all internal self-tests and is ready for deployment.

The defibrillator must provide a visual and audible indicator when: - An error is detected during self-testing - The defibrillator has been stored outside of the prescribed temperature range - The battery is low, depleted, or missing

The defibrillator must test for the usability of the pads using gel moisture as a measure of impedance on a daily basis.

Defibrillator Energy Source Specifications:

Specifications

Standard Battery

The defibrillator must utilize a lithium battery technology that is disposable and recyclable.

The primary battery must provide an operating capacity of at least 200 full energy shocks or at least 4 hours of “On” time or a Standby time that is typically at least 4 years (3 years minimum).

Defibrillation Pads and Cable Specifications:

Specifications

Defibrillation pads must be pre-connected into the defibrillator.

Comprehensive placement icons must appear on each defibrillation pad.

The adhesive pads should be compatible to patients of any age or weight, eliminating the need for a separate defibrillation pads for infant/child applications.

An infant/child key should be available to automatically reduce the defibrillation energy to 50 joules for pediatric application.

Data Collection and Review Specifications:

Specifications

The defibrillator must provide the means to collect and store up to 15 minutes of ECG data and unlimited Event data.

Software that is PC compatible to download and review event data must be available (optional).