Unique Differentiators

Brand: Philips

Model: Efficia DFM100

Physical Specifications:

- 7 inch diagonal LCD TFT color screen and displays 2-3 physiologic waveforms.

- With display resolution of 800 x 480 pixels (VGA).

- Weights 5.66 kg without battery.

- With ready for use indicator to confirm that the device has passed its hourly, daily, and weekly automated self-tests.

- With patients category button to quickly allow clinicians to switch between patient categories

- With smart select knob for easy menu navigation

- With sync button that lights up when in use

- Battery:

- Lithium-ion

- At least 2.5 hours monitoring on full battery

- At least 100 full energy charge/shock cycles on a full battery

- External Paddles:

- External adult paddles with removable adaptors to reveal pediatric paddles

- Has patient contact indicator located in the paddle that lights up to signify good contact with skin

- With remote charge and shock buttons present on the paddles

- Pads:

- Can use multi-function disposable pads for cardioversion, monitoring, AED, or pacing.

- Single pad set for both adult and infant/child patients.

- Built in thermal printer that is compatible with 50mm thermal paper

- Can be configured to automatically print when charging, shocking, alarm events or marking events.

- Code summary can be printed after therapy with real time stamps.

- With patient vital signs trending up to 8 hours.

- With USB port to download event summary

- Meets Ingress Protection level IP44

Operational Specifications:

- Monitor mode:

- Can perform 3 or 5 lead ECG monitoring, with leads off sensing

- Capable of NIBP Monitoring (optional)

- Capable of SPO2 monitoring  (optional)

- Capable of CO2 monitoring using mainstream or sidestream capnography (optional)

- Manual defibrillation:

- Uses low energy FDA-approved biphasic truncated exponential waveform technology with impedance compensation.

- Energy selection up to 200 joules maximum via rotating knob

- Less than 5 secs charging time to recommended adult dose of 150J with a new fully-charged battery

- Capable of synchronized cardioversion

- AED mode: (optional)

- AED mode with voice prompts and CPR coaching

- Validated for adult and infant/child use

- Optimized charging and CPR timing that allows quick shock after CPR

- Continuously monitor patients for shockable rhythms and advises shocks during AED

- AED Energy Profile: 150 Joules for Adult/50 J for Infant/Child nominal into a 50 ohm test load

- Configurable escalating energy on AED mode.

- Transcutaneous pacing: (optional)

- Capable of fixed and demand pacing

- Pace pulse duration should be configurable to 20 or 40ms