Sunray Medical Apparatus Co., Ltd--------Version: V4.1.—2013.05

| SRF618E (FETAL HEART RATE DOPPLER ) SPECIFICATIONS |  |
| --- | --- |
| Figure |  |
| Characteristics | ◆ FHR detection from 10-week pregnancy up to delivery. ◆ ★3 display modes: Transient display mode, average display mode, manual counting model. ◆ ★FHR sound intensity display synchronously ◆ ★Waterproof USB probe for easy replacement. ◆ ★Auto-gain control technology to assure accurate FHR detection. ◆ Water-drip key tone ◆ Build-in speaker. ◆ LED backlight display. ◆ Intellectual cancellation of erroneous FHR. ◆ 16-chord alarming on FHR value over-limitation. ◆ Function on volume adjustment and button click sound. ◆ Three Alkaline AA batteries for 10 hours continuous working. Low battery alarming and real-time battery status indicator. ◆ Auto power-off to save battery if no FHR signal is received within 2 minutes. |
|  | Power Supply: Battery for standard: |
|  | Battery type 3pcs Alkaline AA Power 4.5V |
|  | Continuous working ≥10hours |
| Technical | Battery for option: |
| Parameter | Battery type 3pcs Ni-MH rechargeable batteriesPower 3.6V |
|  | Recharging time ≥5hours Battery volume 1800 mAH |

Continuous working ≥12hours Indicator:

Page 1 of 2

Sunray Medical Apparatus Co., Ltd--------Version: V4.1.—2013.05

| SRF618E (FETAL HEART RATE DOPPLER ) SPECIFICATIONS |  |
| --- | --- |
|  | Volume indication bar Volume station indicator Alarming indicator |
|  | Battery volume indicator Battery type indicator FHR display |
|  | Auto average FHR detection Manual FHR detection Operating Environment: Temperature +5~+40℃ Relative humidity ≤80% |
|  | Atmospheric pressure 86~106KPa. Storage Environment: Temperature –10~+40℃ Relative humidity ≤80% |
|  | Remark: Free from corrosive gases, well ventilated. Dimension and Weight: Dimension 130*78*28mm Gross Weight 0.5kg Display: Digital LCD Fetal Heart Rate (FHR) Measuring: Transient Measuring method Ultrasonic Doppler Measuring mode Average Manual |
|  | Maximal measuring range 30~240BPM Resolution 1BPM |
|  | Integrative sensitivity >100dBUltrasonic operating frequency2.5MHz |
|  | Ultrasonic output intensity ≤5mW/cm2 Audio output power (max) <1W Waterproof probe Yes Alarm: Function Audible and visual alarming with song for abnormal FHR |
|  | Upper limit 160BPM Lower limit 120BPM |
|  | Alarm triggering time <3s (from out-of-limit of FHR to start of alarm.) Auto Power Off: Power off automatically with no signals of FHR received within 2 minutes. |
| Standard Configure | Model Quantity(pcs) Remark Fetal Heart Rate Doppler Detector1 With the cylinder-like water-proof probe |
|  | Alkaline AA Size) Battery 3 |
|  | Packing List 1 |
|  | User’s Manual 1 |
| For Option | External Ni-MH Rechargeable Battery3 |
|  | External Charger 1 Input AC100V~240V |
|  | Canvas Bag 1 |
|  | Earphone 1 |

NOTE: The specifications subject to change without prior notice.

Page 2 of 2