Product			:  Wireless Treadmill Stresstest Machine

Brand			: Philips

Country of Origin	: US

GENERAL SPECIFICATION OF THE SYSTEM

- Must have at least 100 different user profiles

- Able to program up to 100 exercise protocols including Bruce, Modified Bruce, NIH Protocols

- Able to program pharmacological protocols such as Stress Dobutamine for patients unable to utilize the stress test exercise machine

- With standard 5 phases of patient exercise test: Prepare, Pre-exercise, Exercise, Recovery and Report (Post-recovery)

- Must have an automatic 12-lead ECG system and prompts for blood pressure measurement

- Must use DXL and CAIg-STR Exercise ECG Analysis Algorithm

- Must have multi-final report formats with custom report sequence capabilities

- Should have ST Maps and able to present and highlight the ST level & slope status of the patient through zoomed ECG waveforms showing ST level & slope.

- Must display hemodynamic changes like target/maximum/current heart rate, current and previous BP, SpO2, double-product and METS.

- Must have beat by beat review mode, post-test playback allows full disclosure of entire test

- Can interface with the following: Trackmaster TMX425 series threadmill, Ergoline and lode ergometers, Analog and TTL output for stress-echo and NIBP sync, and SunTech Tango NIBP/SpO2 monitor.

Treadmill

- User weight capacity: maximum of 500 lbs / 227 Kg

- Must have a stop / emergency safety feature

- Must have safety tubular handlebars and handrails installed

- Must have cable that is able to connect to CPU

- Speed and inclination must be manually adjustable:

  - Speed range should be adjustable from 0.1-12mph (.16-19kph) or more

  - Inclination must be adjustable from 0-25% or more

Transmitter

- Must have a 10-wire lead ECG cable, AAMI or IEC color coding, lead replacement must be standard

- Must be wireless and uses commercially available single AA alkaline battery

- Must be compact and lightweight, 900g or less

- Must have a battery life indicator/power status

- Must have lead map with lead connection status LEDs

- Must have signal strength indicator

- System must be able to pair with more than one transmitter for patients in queue

Monitor

- All-In-One System (Monitor and CPU must be one system)

- At least 23-inch LED, colored, flat panel, touchscreen

- All instructions and display read out should be in English

CPU

- All-In-One Monitor and CPU must be one system

- At least have an Intel Quad Core i3

- At least Microsoft Windows 8.1 Pro license

- Must have at least 4GB RAM

- Must have Storage Hard drive of at least 500GB

- Must have a keyboard, optical mouse, Bluetooth

- Must be able to be upgraded up to the latest Microsoft version

- Minimum 4USB Ports

- Optical DVD/CD Drive

UPS (Third Party)

- Voltage: 240 VAC at 50/60 Hz

- Life / standing time when not connected to main power switch: must be at least 30 minutes

Printer (Third Party)

- Must be a high-speed Laser jet, 30 ppm minimum

- Must be able to print in black and white

- Print resolution: Up to 1200 by 1200 dpi

- Must have a print server with the following features

  - Must have a processor of at least 540 MHz

  - Must have memory of at least 128 MB (upgradeable to 640 MB – free of service)

  - Must have a security lock slot

  - Must have a Hi-Speed USB 2.0 port

Trolley

- Manufacturer OEM Trolley, customized for stress test system

- Must have shelves to house CPU, monitor, keyboard, printer, and UPS

- VESA type monitor mount, swivel or fixed

- Built-in pull-out drawer

- Powder coated steel finish

- Must have lockable caster wheels

Software

- Software must be able to differentiate between noise and signal in each lead while preserving diagnostic-quality ECG waveforms

- Must be PC based with minimum Microsoft Windows operating system (upgradeable).

- Must have wireless ECG acquisition to eliminate noise and other problems associated with cable sag and movements (transmitter)