## Sheet1
Name | Video Laryngoscope
Model | VS-10H
Function | Designed to be used to facilitate the openning of airway
Dimensions | 191(H) x 92 (W) x 112 (D) mm
Weight | About 0.25kg (including the battery)
Power Adapter | Input voltage: AC 100-240V 50/60Hz.
Input power: 25VA.
DC output voltage: 5V 2A.
Battery | Built-in battery: 3.6V 3400mAh.
Continuous operation duration of the lithium battery: The device can operate for at least 4 hours if a brand new fully-charged battery is used.
Lithium battery charge time: not longer than 4 hours
Lithium battery charge mode: The lithium battery is charged through the USB cable
Rotation Angle of Display Screen | Maximum vertical rotation angle: 140°±10°
Maximum horizontal rotation angle: 270°±10°
Spatial/Image Resolution | Not smaller than 6.5 lp/mm
Color Restoration Capability | Not smaller than level 4
Light Source | LED light source brightness can be adjusted to adapt to different use environments.
Display Brightness | Featuring automatic white balance adjustment, high resolution, clear image, adjustable display screen brightness
Screen | 3.5 inches touchscreen with resolution of 640x960 pixels
The color temperature of the LCD screen at least 5000 K
Depth of Field | 10-80mm
Field Angle | Not smaller than 60°
Viewing Angle | 0°±10°
Illumination Intensity | Not smaller than 600lx
Light Source Color Temperature | Not smaller than 5000K
Circular Arc Radius of Handle Insertion Part | Circular arc radius: 75mm; deviation: ±5%
Water Proof Level | IP66;
Phote Record | Photo can be recorded in JPEG format, up to 10,000 photos can be recorded with Snapshot function
Vidoe Record | Video can be recorded in MPEG-4 format, up to 2 hours video can be recorded
Media Storage | Not less than 5GB
Data Port | HDMI, USB Tpye C and Wi-Fi are available
Data Transmission | Both wired and wireless file transmission are supported, providing better teaching experience
Teaching Function | When the device is connected to the HDMI cable, the video can be displayed on an output screen in real-time
Anti-Fog | Immediate anti-fog is supported with no need for heating the anti-fogging coating on the laryngoscope blade
Blades | Single use blade.
Macintosh blade M1,M2,M3 and M4 available, with guiding channel
M1:106(L)x18(W)x33mm; M2: 115(L)x22(H)x33(W)mm; M3:126(L)x27(W)x38(H)mm; M4: 140(L)x29(W)x41(H); M3D: 126(L)x24(W)x42(H) mm (L: ±5mm, W: ±3mm, H: ±3mm)
The focus length for adult laryngoscope is 10-80mm
The focus length for pediatric laryngoscope is 10-80mm
Special Blade | Special adult blade M3D for difficult intubation
Frequency for WiFi | 2.4GHz
Certificate | ISO 13485, FDA, CE